“EVO-SSL-670-5” Protocol Description

Contents

1.	Contents	2

2.	Description of the UKKO_NS_1064_Interface_board protocol	3

3.	Command format:	3

4.	Common commands	4

4.1.	Get protocol version	4

4.2.	Get hash	4

4.3.	Factory reset request	4

4.4.	Switch to Bootloader request	4

5.	User commands	5

5.1.	Request to get status	5

5.2.	Request to set emission on/off state	5

5.3.	Request to set mode	5

5.4.	Request to set mode	6

5.5.	Request to set PSU permission	6

5.6.	Request to clear errors	6

5.7.	Request to set quantron pulse frequency	7

5.8.	Request set pulse duration	7

6.	Admin commands	8

6.1.	Status request	8

6.2.	Request set configuration	8

6.3.	Request get configuration	8

6.4.	Request save configuration	9

Description of the UKKO_NS_1064_Interface_board protocol

Supported connection types: Ethernet, USB.

Protocol provides simple text commands and responses.

On board TCP server parameters:

IP address type: Static

IP address: some_ip (192.168.26.220)

Server port: 7878

COM port parameters:

Baud rate: 115200 kbps

Parity: none

Stop bits: 1

Data bits: 8

Command format:

All commands and responses are a simple text string that has the following structure:

<message_header> <sub_sys_type> <id> <param1> <param2> <…> <CR>

Where:

- <message_header> - message header that contains command keyword

- <sub_sys_type> - target subsystem addressed by the command

- <id> - subsys address

NOTE: For this laser configuration, the <id> is always 0

- <paramN> - a bunch of parameters corresponding to command

- <CR> - carriage return symbol

Common commands

Protocol version name: " EVO-UC-NS-775_180326".

Get protocol version

“lgvers\r”

Response:

“lrvers <protocol version name>\r”

Where:

- <protocol version name> - protocol version name (See. Chapter 1)

Get hash

“lghash\r”

Response:

“lrhash <hash text>\r”

Where:

- < hash text> - information for developers.

Factory reset request

“lsfactrst\r”

Response:

“lrfactrst OK\r”

ERRORE response:

“lrfactrst ERR\r”

Switch to Bootloader request

“lsbootloader\r”

Response:

“lrbootloader ON\r”

NOTE:

- In a real working hardware, the response command will most likely not come

User commands

Request to get status

"lgstatus usr <id>\r"

Response:

"lrstatus usr <id> <interlock><chiller_fail> <overheat> <ld_temp_not_ready> <seed_state> <PSU_permission> <PSU_state> <quant_mode> <pocels_div> <temp1> <temp2> <tec_temp1> <tec_temp2> <tec_temp3> <tec_temp4>\r\n"

ERROR response:

“lrstatus usr <id> ERR\r”

Where:

- <id> - device id (int), 0 if device is single

- <interlock> - interlock state (int), 0 – ok, 1 – alarm

- <chiller_fail> - chiller fail (int), 0 – ok, 1 – alarm

- <overheat> - overheat state (int), 0 – ok, 1 – alarm

- <ld_temp_not_ready> - Laser Diode temperature not ready (int), 0 – ok, 1 – alarm

- <seed_state> - seed emission state (int), 0 – off, 1 – on

- <PSU_permission> - power supply work permission (int) 0 – off, 1 – on

- <PSU_state> - power supply actual state (int), 0 - off, 1 – on

- <output_state> - output emission state (int), 0 – off, 1 – on

- <quant_mode> - quantron sync mode (int) 0 – Internal, 1 – External

- <pocels_div> - pocels divider (int) 1-254

- <tempN> - N temperature (float) 0.0 – 200.0 C

- <tec_tempN> - tec N temperature (float) 0.0 – 200.0 C

Request to set emission on/off state

“lsonoff usr <id> <value>\r”

Response:

“lronoff usr <id> <value>\r”

ERROR Response:

“lronoff usr <id> ERR\r”

Where:

- <id> - device id (int), 0 if device is single

- <value> - emission state value, (int) 0 – off, 1 – on

Request to set mode

"lsmode usr <id> <value>\r"

Response:

“lrmode usr <id> <value>\r"

ERROR response:

“lrmode usr <id> ERR\r”

Where:

- <id> - device id, (int); always 0.

- <value> - quantron sync mode (int) 0 – Internal, 1 – External

Request to set mode

"lsdiv usr <id> <value>\r"

Response:

“lrdiv usr <id> <value>\r"

ERROR response:

“lrdiv usr <id> ERR\r”

Where:

- <id> - device id, (int); always 0.

- <value> - pocels puls divider (int) 1 – 254

Request to set PSU permission

"lspsupermission usr <id> <value>\r"

Response:

“lrpsupermission usr <id> <value>\r"

ERROR response:

“lrpsupermission usr <id> ERR\r”

Where:

- <id> - device id, (int); always 0.

- <value> - permission value (int), 0 – blocked, 1 - allowed;

Request to clear errors

"lserrclr usr <id>\r"

Response:

“lrerrclr usr <id>\r"

ERROR response:

“lrerrclr usr <id> ERR\r”

Where:

- <id> - device id, (int); always 0

Request to set quantron pulse frequency

"lsfreq usr <id> <value>\r"

Response:

“lrfreq usr <id> <value>\r"

ERROR response:

“lrfreq usr <id> ERR\r”

Where:

- <id> - device id, (int); always 0.

- <value> - frequency value (int), 1 – 100 Hz

Request set pulse duration

"lsdura usr <id> <value>\r"

Response:

"lsdura usr <id> <value>\r"

ERROR response:

"lsdura usr <id> ERR\r"

Where:

- <id> - device id (int), 0 if device is single

- <value> - pulse duration (float), range of values: 50 – 300 us

Admin commands

Status request

“lgstatus admn <id>\r”

Response:

“lrstatus admn <id> <power_led> <PSU_led> <orange_led> <red_led> <err_code>\r”

ERROR Response:

“lrstatus admn <id> ERR\r”

Where:

- <id> - device id (int), 0 if device is single

- <power_led> - green power led state, (int) 0 – off, 1 - on

- <PSU_led> - PSU green led state, (int) 0 – off, 1 - on

- <orange_led> - emission orange led state, (int) 0 – off, 1 - on

- <red_led> - error red led state, (int) 0 – off, 1 – on

- <err_code> - error code, (int)

Request set configuration

"lsconf admn <id> <tec_state0> <tec_temp0> <tec_state1> <tec_temp1> <tec_state2> <tec_temp2> <tec_state3> <tec_temp3> <hpld_curr> <tec_max0> <tec_max1> <tec_max2> <tec_max3> <temp_max0> <temp_max1>\r"

Response:

"lsconf admn <id> <tec_state0> <tec_temp0> <tec_state1> <tec_temp1> <tec_state2> <tec_temp2> <tec_state3> <tec_temp3> <hpld_curr> <tec_max0> <tec_max1> <tec_max2> <tec_max3> <temp_max0> <temp_max1>\r"

ERROR response:

"lrconf admn <id> ERR\r"

Where:

- <id> - device id (int), 0 if device is single

- <tec_stateN> - N TEC state value (int), 0 – off, 1 – on

- <tec_tempN> - N TEC temperature setpoint (float), 10.00 – 200.00 C

- <hpld_curr> - hpld-1000 stored current value (float), 0.00 – 26.00 A

- <tec_maxN> - N TEC max temperature value (float), 0.00 – 200.00 C

- <temp_maxN> - N temperature max value (float), 0.00 – 200.00 C

Request get configuration

"lgconf admn <id>\r"

Response:

"lsconf admn <id> <tec_state0> <tec_temp0> <tec_state1> <tec_temp1> <tec_state2> <tec_temp2> <tec_state3> <tec_temp3> <hpld_curr> <tec_max0> <tec_max1> <tec_max2> <tec_max3> <temp_max0> <temp_max1>\r"

ERROR response:

"lrconf admn <id> ERR\r"

Where:

- <id> - device id (int), 0 if device is single

- <tec_stateN> - N TEC state value (int), 0 – off, 1 – on

- <tec_tempN> - N TEC temperature setpoint (float), 10.00 – 200.00 C

- <hpld_curr> - hpld-1000 stored current value (float), 0.00 – 26.00 A

- <tec_maxN> - N TEC max temperature value (float), 0.00 – 200.00 C

- <temp_maxN> - N temperature max value (float), 0.00 – 200.00 C

Request save configuration

"lssave admn <id>\r"

Response:

"lrsave admn <id> OK\r"

ERROR response:

"lrsave admn <id> ERR\r"

Where:

- <id> - device id (int), 0 if device is single